PART 2 – Supply Requirements

#### Contents

- List of Goods and Delivery Schedule	49

- List of Related Services and Completion Schedule	51

- Technical Specifications	55

- Drawings	83

- Inspections and Tests	84

Notes for Preparing the Schedule of Requirements

The Schedule of Requirements shall be included in the bidding documents by the Purchaser, and shall cover, at a minimum, a description of the goods and services to be supplied and the delivery schedule.

The objective of the Schedule of Requirements is to provide sufficient information to enable bidders to prepare their bids efficiently and accurately, in particular, the Price Schedule, for which a form is provided in Section IV. In addition, the Schedule of Requirements, together with the Price Schedule, should serve as a basis in the event of quantity variation at the time of award of contract pursuant to ITB 41.

The date or period for delivery should be carefully specified, taking into account (a) the implications of delivery terms stipulated in the Instructions to Bidders pursuant to the Incoterms rules (i.e., EXW, or CIP, FOB, FCA terms—that “delivery” takes place when goods are delivered to the carriers), and (b) the date prescribed herein from which the Purchaser’s delivery obligations start (i.e., notice of award, contract signature, opening or confirmation of the letter of credit).

- List of Goods and Delivery Schedule

All items shall to be delivered to each Project Site:

- National hospital under MH KR. KZ,  720040 Bishkek,  ul. Togoloka Moldo,1

- Zhalal-Abadskaya  clinical hospital. KZ,  720901 Zhalal-Abad, ul.Pushkina 91

Delivery shall be done during not more than 120 calendar days upon signing the Contract

| Line item # | Description of Goods | Quantity For each project site | Physical unit | Final (project Site) Destination : | Delivery (as per Incoterms) Date |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Earliest Delivery Date | Latest Delivery date | Bidder’s offered. Delivery date (to be provided by the bidder) |
| I | Module I |  |  | - National hospital under MH KR. KZ, 720040 Bishkek, ul. Togoloka Moldo,1 - Zhalal-Abadskaya clinical hospital. KZ, 720901 Zhalal-Abad, ul.Pushkina 91 | As soon as possible | 120 days upon signing the Contract |  |
| 1. | Vehicle (see details in Sectiion 3.Technical Specifications) | 1 | pcs |  |  |  |  |
| 2. | Medical equipment |  |  |  |  |  |  |
| 2.1. | Ultraviolet bactericidal air recirculator | Not less than 2 | pcs |  |  |  |  |
| 2.2. | Digital low-dose photofluorographic unit | 1 | pcs |  |  |  |  |
| II | Module 2 |  |  | - National hospital under MH KR. KZ, 720040 Bishkek, ul. Togoloka Moldo,1 - Zhalal-Abadskaya clinical hospital. KZ, 720901 Zhalal-Abad, ul.Pushkina 91 | As soon as possible | 120 days upon signing the Contract |  |
| 1. | Vehicle (see details in 3.Technical Specifications) | 1 | pcs |  |  |  |  |
| 2. | Medical equipment |  |  |  |  |  |  |
| 2.1 | Dental unit | Not less than 1 | pcs |  |  |  |  |
| 2.2 | X-ray dental device. | Not less than 1 | pcs |  |  |  |  |
| 2.3 | The intraoral radiology dental | Not less than 1 | pcs |  |  |  |  |
| 2.4 | Glasperlen sterilizer | Not less than 1 | pcs |  |  |  |  |
| 2.5 | Ultrasonic bath | Not less than 1 | pcs |  |  |  |  |
| 2.6 | Dry-burning Cabinet | Not less than 1 | pcs |  |  |  |  |
| 2.7 | UV chamber for preserving sterility | Not less than 1 | pcs |  |  |  |  |
| 2.8 | Container for disinfection and pre-sterilization treatment of medical devices | Not less than 2 | pcs |  |  |  |  |
| 2.9 | Ultraviolet bactericidal air recirculator | Not less than 2 | pcs |  |  |  |  |
| 2.10 | Tester the viability of the pulp | Not less than 1 | pcs |  |  |  |  |
| 2.11 | Apex locator | Not less than 1 | pcs |  |  |  |  |
| 2.12 | The electrocoagulator | Not less than 1 | pcs |  |  |  |  |
| 2.13 | Led device | Not less than 1 | pcs |  |  |  |  |
| 2.14 | Portable autoclave | Not less than 1 | pcs |  |  |  |  |
| 2.15 | A set of tools | Not less than 1 | pcs |  |  |  |  |
| 2.16 | General practitioner bag with medical diagnostic kits | Not less than 1 | pcs |  |  |  |  |
| III | Module 3 |  |  | - National hospital under MH KR. KZ, 720040 Bishkek, ul. Togoloka Moldo,1 - Zhalal-Abadskaya clinical hospital. KZ, 720901 Zhalal-Abad, ul.Pushkina 91 | As soon as possible | 120 days upon signing the Contract |  |
| 1 | Vehicle (see details in 3.Technical Specifications) | 1 | pcs. |  |  |  |  |
| 2 | Medical equipment |  |  |  |  |  |  |
| 2.1 | Gynecological examination chair | At least 1 | pcs. |  |  |  |  |
| 2.2 | Wall-mounted lighting fixture | At least 2 | pcs. |  |  |  |  |
| 2.3 | Colposcope on a console support | At least 1 | pcs. |  |  |  |  |
| 2.4 | Microscope | At least 1 | pcs. |  |  |  |  |
| 2.5 | Portable ultrasonic system with accessories | At least 1 | pcs. |  |  |  |  |
| 2.6 | Manual artificial respirating unit | At least 1 | pcs. |  |  |  |  |
| 2.7 | A general kit for delivery of first aid (pcs) | At least 1 | pcs. |  |  |  |  |
| 2.8 | Electrocardiograph | 1 | pcs. |  |  |  |  |
| 2.9 | Diagnostic spirometer | At least1 | pcs. |  |  |  |  |
| 2.10 | Stethoscope | At least 2 | pcs. |  |  |  |  |
| 2.11 | Cardiological stethoscope | At least 1 | pcs. |  |  |  |  |
| 2.12 | Wall-mounted sphygmomanometer with a dial | At least 3 | pcs. |  |  |  |  |
| 2.13 | Height meter | At least 1 | pcs. |  |  |  |  |
| 2.14 | Medical balance | At least 1 | pcs. |  |  |  |  |
| 2.15 | Pedobarometer (pcs) | At least 1 | pcs. |  |  |  |  |
| 2.16 | Biochemical portable tester of blood | At least 1 | pcs. |  |  |  |  |
| 2.17 | Glucometer | At least 1 | pcs. |  |  |  |  |
| 2.18 | Portable urine analyser | At least 1 | pcs. |  |  |  |  |
| 2.19. | Stand | At least 2 | pcs. |  |  |  |  |
| 2.20 | Container for disinfection and presterilising treatment of medical devices | At least 4 | pcs. |  |  |  |  |
| 2.21 | Ultraviolet bactericidal air recirculator | At least 3 | pcs. |  |  |  |  |
| 2.22 | Blood analyzer hematological | At least 1 | pcs. |  |  |  |  |
| IV | Module 4 |  |  | - National hospital under MH KR. KZ, 720040 Bishkek, ul. Togoloka Moldo,1 - Zhalal-Abadskaya clinical hospital. KZ, 720901 Zhalal-Abad, ul.Pushkina 91 | As soon as possible | 120 days upon signing the Contract |  |
| 1. | Vehicle (see details in 3.Technical Specifications) | 1 | pcs. |  |  |  |  |
| 2. | Medical equipment: |  |  |  |  |  |  |
| 2.1. | Ophthalmoscope with accessories (pcs) | At least 1 | Pcs. |  |  |  |  |
| 2.2. | Otoscope with accessories | At least 1 | pcs. |  |  |  |  |
| 2.3. | A set of test spectacle lenses with a frame | At least 1 | pcs. |  |  |  |  |
| 2.4. | Light for visual charts | At least 1 |  |  |  |  |  |
| 2.5. | Air-puff tonometer for ophthalmotonous pressure measurement | At least 1 | pcs. |  |  |  |  |
| 2.6. | Manual portable slit lamp | 1 | pcs. |  |  |  |  |
| 2.7. | Stethoscope | At least 2 | pcs. |  |  |  |  |
| 2.8. | Wall-mounted sphygmomanometer with a dial | At least 3 | pcs. |  |  |  |  |
| 2.9. | Ultraviolet bactericidal air recirculator | 4 | pcs. |  |  |  |  |

### 2.List of Related Services and Completion Schedule

NOT APPLICABLE

| Service | Description of Service | Quantity1 | Physical Unit | Place where Services shall be performed | Final Completion Date(s) of Services |
| --- | --- | --- | --- | --- | --- |
| [insert | [insert description of Related Services] | [insert quantity of | [insert physical | [insert name of the | [insert |
| Service |  | items to be | unit for the items] | Place] | required |
| No] |  | supplied] |  |  | Completion |
|  |  |  |  |  | Date(s)] |

### 3. Technical Specifications

# Mobile modular complex for providing first aid and primary medical-sanitary care to residents of remote and hard-to-reach districts and settlements of the Kirgiz Republic “Karavan zdoroviya”(“Health Caravan”).

Delivery period requirement – not more than120 calendar days upon signing of the Contract

Module 1

| No. | Item | Item content (value) |
| --- | --- | --- |
| I. | Module 1 | on a vehicle chassis |
| 1. | General characteristics of Module 1 |  |
| 1.1 | Purpose | For delivering medical aid during parking (stops) |
| 1.2 | Compliance with the requirements of TR CU 018/2011 of the Technical Regulations of the Customs Union. On the safety of wheeled vehicles. | Availability of confirmation in the form of a randomly completed application, certified by the seal of the company. |
| 1.3 | Category of vehicle | N3G |
| 1.4 | Wheel arrangement of the vehicle | 6х6 |
| 1.5 | Axles quantity | Not less than 3 |
| 1.6 | Cabin type | Over the engine location |
| 1.7 | Number of seats in the cabin | Not less than 3 |
| 1.8 | Engine type | Diesel, four-stroke |
| 1.8.1 | Emission class | Euro-5 |
| 1.8.2 | Cubic capacity of the engine, l | Not less than 6,7 |
| 1.9 | Van type | Multi-section insulated van consisting of sandwich panels with fixed partitions mounted on the vehicle chassis |
| 1.10 | Van dimensions |  |
| 1.10.1 | Length, m | Not less than 5,5 |
| 1.10.2 | Width, m | Not less than 2,4 |
| 1.10.3 | Hight, m | Not less than 2,4 |
| 1.11 | The composition of the exterior sandwich panels | - clad steel coated with white powder enamel or white plastic reinforced with fiberglass, resistant to disinfectant treatment; -plywood; expanded polystyrene; -plywood; - white plastic reinforced with fiberglass, resistant to disinfectant treatment |
| 1.12 | The composition of the inner sandwich panels | - white plastic reinforced with fiberglass, resistant to disinfectant treatment; -plywood; expanded polystyrene; -plywood; - white plastic reinforced with fiberglass, resistant to disinfectant treatment |
| 1.13 | The width of the external sandwich panels, mm | Not less than 60 |
| 1.14 | The width of the inner sandwich panels, mm | Not less than 40 |
| 1.15 | The elements embedded in the sandwich panels for installation of internal and external equipment | Availability |
| 1.16 | Compartments |  |
| 1.16.1 | Medical treatment compartment | with a wing door |
| 1.16.2 | Sanitary compartment | with a wing door |
| 1.16.3 | Generator compartment | with a wing door |
| 1.16.4 | Climate control compartment | with a wing door |
| 1.17 | Window | Tinted double-pane glass with light-protective shutters |
| 1.17.1 | Glazing area | Not less than 700×500 |
| 1.18 | Doors | of sandwich panels with locking devices, holding devices in closed and open position |
| 1.18.1 | The width of the doors sandwich panels, mm | Not less than 40 |
| 1.18.2 | The doors dimensions (HxW), mm | Not less than 1800×600 |
| 1.19 | Floor | unibody construction of sandwich panels laid with linoleum |
| 1.19.1 | The width of the floor sandwich panel, mm | Not less than 80 |
| 1.20 | Accessories (hinges, locks), exterior and interior trim of the van, doors, door openings, joints of the sandwich panels in various plane surfaces | Stainless steel, aluminum extrusion |
| 1.21 | Entry portal in the corridor | A structure made of aluminium shape with double packing |
| 1.22 | The box beneath the van bottom | Metal welded box with locking device |
| 1.23 | External power supply | Connection to single-phase power grid 220 Volt, m |
| 1.23.1 | The cable length for connection to single-phase power grid 220 Volt, m | Not less than 30 |
| 1.24 | Protective grounding | Availability |
| 1.25 | Autonomous power supply | Gasoline single-phase power supply station, 220 Volt |
| 1.25.1 | Petrol-powered electric generation station capacity, kW | Not less than 9 |
| 1.25.2 | The tank volume of the petrol-powered electric generation station, l | 10 |
| 1.26 | Autonomous power supply from the on-board power supply system of the vehicle | Via a multifunctional voltage converter with 24 V DC on-board vehicle network to 220 V AC |
| 1.26.1 | Power output of multifunction voltage converter, kW | Not less than 2 |
| 1.27 | Electrical equipment | Input electrical panel with switches for external and autonomous power supply, input voltage control relay, introductory electric automatic circuit breaker and incoming line section of safety shutdown, electrical consumer lines with differential circuit breakers, connectors for external and autonomous power supply. Electrical accessories (sockets, switches) in each compartment |
| 1.28 | The number of individual lines of electrical consumers with differential automatic circuit breakers | Not less than 6 |
| 1.29 | Conditioning system | Split system |
| 1.29.1 | External unit of a split-system | Located in the climate control compartment |
| 1.29.2 | Internal unit of a split system | Located in the medical treatment compartment |
| 1.29.3 | Conditioning system operation modes | cooling, ventilation, heating |
| 1.30 | Ventilation system |  |
| 1.30.1 | Ventilation system of the medical treatment compartment | Supply air filtering and ventilation unit with prepared air distribution |
| 1.30.2 | Ventilation system of the generator compartment | Forced supply and exhaust ventilation system with the capacity sufficient for normal operation of a gasoline power station |
| 1.30.3 | Ventilation system of the climate control compartment | Forced supply and exhaust ventilation system by external unit of the split-system |
| 1.31 | Heating system |  |
| 1.31.1 | Convector heater, pcs. | Not less than 2 |
| 1.31.1.1 | Capacity of the convector heater in the medical treatment compartment, kW | Not less than 1 |
| 1.31.1.2 | Capacity of the convector heater in the sanitary compartment, kW | Not less than 0,5 |
| 1.31.2 | Air heat curtain | Above the entry opening in the corridor |
| 1.31.2.1 | Capacity of the air curtain, kW | Not less than 2 |
| 1.31.3 | Autonomous diesel fuel air heater | with the capacity to operate on the move and during parking |
| 1.31.3.1 | Capacity of the autonomous diesel fuel air heater, kW | Not less than 4 |
| 1.32 | The basic lighting system | light emitting diode lighting units |
| 1.33 | The emergency lighting system | from the on-board power supply system of the vehicle, 24 V |
| 1.33.1 | Number of emergency lighting unit in each compartment, pcs. | Not less than 1 |
| 1.34 | Fire-fighting appliances | Fire alarm system with the smoke sensors in each compartment; powder-type fire extinguisher |
| 1.34.1 | Capacity of the powder-type fire extinguisher, l | Not less than 2 |
| 1.35 | Sanitary equipment | Availability |
| 1.35.1. | The sink with a tap and a pump for feeding the prepared and heated water in the sanitary compartment | Availability |
| 1.35.2 | Cassette type biotoilet in the sanitary compartment | Availability |
| 1.36 | Furniture | Constructed of panels with laminated working surfaces robust to disinfectant treatment |
| 1.36.1 | Suspended double door shelf in the medical treatment compartment | Availability |
| 1.36.2 | Table in the medical treatment compartment | Availability |
| 1.36.3 | Castor chair with a gas-lift unit in the medical treatment compartment, pcs | Not less than 2 |
| 1.36.4 | Stool in the corridor | Not less than 2 |
| 1.37 | Utility equipment of the compartments | Availability |
| 1.37.1 | Soap dispensers in the sanitary compartment | Availability |
| 1.37.2 | Towel rail in the sanitary compartment | Availability |
| 1.37.3 | Bins, pcs | Not less than 2 |
| 1.37.4 | Clothes hooks, pcs | Not less than 2 |
| 1.37.5 | Container for clean shoe covers, pcs | Not less than 1 |
| 1.38.6 | Container for discarded shoe covers, pcs | Not less than 1 |
| 1.39 | System for fixing medical equipment and other equipment in transport position | Availability |
| 1.40 | Access entry | including the stairs integrated in the van bottom and additional attachable stairs |
| 1.41 | Dashboard camera, pcs. | 1 |
| 1.41.1 | Function | Video and audio record of the situation in front of the vehicle on the move and during parking |
| 1.42 | Subscription telematic server, pcs. | 1 |
| 1.42.1 | Function | Determination and transfer to the telematic server of the location and the movement mode parameters of the vehicle (object) geographical coordinates, speed, course |
| 1.42.2 | GLONASS/GPS receiver | Availability |
| 1.42.3 | GSM Module | Availability |
| 1.43 | Vehicle system of emergency response service, pcs. | 1 |
| 1.44 | Repair kit, pcs. | 1 |
| 1.44.1 | Content of the repair kit | car lifting jack, wheel wrench, warning triangle |
| 2. | Medical equipment |  |
| 2.2 | Ultraviolet bactericidal air recirculator, pcs. | Not less than 2 |
| 2.1.1 | Type | Recirculating |
| 2.1.2 | Number of lamps, pcs. | Not less than 3 |
| 2.1.3 | Lamps service life, h | Not less than 9000 |
| 2.1.4 | Power consumption, kW | Not more than 250 |
| 2.1.5 | Performance m3/h | Not less than 50 |
| 2.1.6 | Sterilization efficacy, % | Not less than 99 |
| 2.1.7 | Weight, kg | Not more than 8 |
| 2.2 | Digital low-dose photofluorographic unit, pcs. | Not less than 1 |
| 2.2.1 | General requirements |  |
| 2.2.1.1 | Non-scanning fluorographic unit, which involves simultaneous shooting of the entire survey area, obtaining a picture directly from a digital detector without intermediate digitization, as well as obtaining a full-format image for one exposure, without intermediate cross-linking of fragments. | Availability |
| 2.2.1.2 | Fluorographic unit must ensure the production of fluorography of the chest organs of the patient with the vertical position of the subject in the straight and lateral projections. | Availability |
| 2.2.2 | Delivery set |  |
| 2.2.2.1 | - A tripod with a built-in (not located outside) x-ray feeding device (RPU) of a monoblock type (emitter). - A tripod with an x-ray image recording device (digital receiver) | Availability |
| 2.2.2.2 | Intercom | Availability |
| 2.2.2.3 | Automated work stations( AWS) for radiologists and radiologists (including the control panel) with software | Availability |
| 2.2.2.4 | Auxiliary devices and accessories | Availability |
| 2.2.3 | Main operational characteristics |  |
| 2.2.3.1 | Supply voltage and frequency | 220V, 50 Hz |
| 2.2.3.2 | Method for getting an image | Direct or indirect |
| 2.2.3.3 | Working field size, mm | Not less than 430 х 430 |
| 2.2.3.4 | Digital x-ray receiver | Availability |
| 2.2.3.5 | Pixel size of the matrix, mcm | Not more than 140 |
| 2.2.3.6 | Direct signal transmission | Availability |
| 2.2.3.7 | The range of exposure time when getting a full-format image, seconds | Not more than 0,03 - 0,06 |
| 2.2.3.8 | Time to display a full-format image on the AWP screen, seconds | Not more than 5 |
| 2.2.3.9 | Number of images per one hour of operation, pcs. | Not less than 60 |
| 2.2.3.10 | Weight of the tripod part, kg | Not more than 250 |
| 2.2.3.11 | The monoblock type RPU is integrated into a tripod and does not have stand alone modules | Availability |
| 2.2.4 | Parameters of the resulting image |  |
| 2.2.4.1 | Full-format image size, pixels | Not less than 3072 х 3072 |
| 2.2.4.2 | Dynamic range measured by the method in accordance with STP 01-22-06, times | Not less than 500 |
| 2.2.4.3 | Digitalization of the x-ray image receiver signal, bit / pixel | Not less than 14 |
| 2.2.5 | Automated work places (AWP) of the radiologist and radiology technician degree |  |
| 2.2.5.1 | Automatic workplace of a radiologist, pcs. | Not less than 1 |
| 2.2.5.2 | Automatic workplace of x-ray laboratory assistant (Touch monoblock), pcs. | Not less than 1 |
| 2.2.5.3 | Number of monitors for the laboratory assistant's AWP, pcs. | Not less than 1 |
| 2.2.5.4 | Diagonal of the laboratory assistant's AWP monitor, inches | Not less than 21 |
| 2.2.5.5 | Diagonals of the doctor's AWP monitor, inches | Not less than 17 |
| 2.2.5.6 | Office printer for printing conclusions, pcs. | Not less than 1 |
| 2.2.5.8 | The ability to connect to the machine additional working places of doctors-radiologists | Availability |
| 2.2.5.9 | Features for the automatic workplace of a radiologist or an automatic workplace of a radiologist: |  |
| 2.2.5.9.1 | The operator interface language is Russian | Availability |
| 2.2.5.10 | Maintaining a patient database: |  |
| 2.2.5.10.1 | Storing the following patient data: Card number, full name, date of birth, gender, group, policy, address, phone number, profession, diagnosis, complaints, comment | Availability |
| 2.2.5.10.2 | Functions for working with the patient database: adding new patients, editing, and deleting patient and study records. Search for the desired patient using any of their personal data, creating a list of patients using various filters | Availability |
| 2.2.5.11 | Maintaining a database of research performed on the device: |  |
| 2.2.5.11.1 | Storage of the following information about the study: exposure parameters (kV, wt, sec), effective radiation dose of the patient, type of study, anatomical area, projection, two protocols of the study with a conclusion, anamnesis and full name of the doctor who created the Protocol, availability of images in the study (for example, a snapshot in the archive, name and type of media) | Availability |
| 2.2.5.11.2 | Functions of working with the research database: adding a new study, deleting a study, adding a description and conclusion of the study, creating lists of studies by various criteria: completed, but not described, performed in the specified time interval | Availability |
| 2.2.5.12 | Work with images: |  |
| 2.2.5.12.1 | Supported image formats: DICOM | Availability |
| 2.2.5.13 | Working with images and related information obtained from the following types of digital equipment: |  |
| 2.2.5.13.1 | digital radiography/radiography equipment (Digital Radiography) | Availability |
| 2.2.5.14 | Viewing and processing images: |  |
| 2.2.5.14.1 | output 16 images simultaneously to the viewing table | Availability |
| 2.2.5.14.2 | output an image to the screen size | Availability |
| 2.2.5.14.3 | rotate the image by a multiple of 90 degrees | Availability |
| 2.2.5.14.4 | image mirroring (vertical/horizontal) | Availability |
| 2.2.5.14.5 | managing image brightness / contrast | Availability |
| 2.2.5.14.6 | viewing images in the negative or positive | Availability |
| 2.2.5.15 | The function of the image post-processing: |  |
| 2.2.5.15.1 | automatic detection of image brightness and contrast | Availability |
| 2.2.5.15.2 | gamma correction | Availability |
| 2.2.5.15.3 | changing the sharpness | Availability |
| 2.2.5.15.4 | noise level | Availability |
| 2.2.5.15.5 | enhancing the contrast of low-contrast objects | Availability |
| 2.2.5.15.6 | the alignment of the image brightness | Availability |
| 2.2.5.16 | Radiometry functions: |  |
| 2.2.5.16.1 | dimensioning | Availability |
| 2.2.5.16.2 | the definition of distances | Availability |
| 2.2.5.16.3 | determination of the angles | Availability |
| 2.2.5.16.4 | the definition of the areas for rectangular and elliptical areas | Availability |
| 2.2.5.16.5 | determining the statistical characteristics of the image and the selected area of interest (minimum, maximum, average brightness, standard deviation of brightness) | Availability |
| 2.2.5.16.6 | estimation of the optical density at a given point | Availability |
| 2.2.5.16.7 | getting an image histogram | Availability |
| 2.2.5.17 | Printing images: |  |
| 2.2.5.17.1 | Print the selected images together with the annotation to A3, A4, and A5 printers. | Availability |
| 2.2.5.17.2 | Printing images to medical thermal printers multi format film cameras defined in the system | Availability |
| 2.2.5.18 | Preparation of research protocols: |  |
| 2.2.5.18.1 | printing research reports to a printer | Availability |
| 2.2.5.18.2 | Generate the following reports: per day, by group, with a description | Availability |
| 2.2.5.18.2 | Printing reports to an office printer | Availability |
| 2.2.5.19 | Exporting research to external media: |  |
| 2.2.5.19.1 | Export research and protocols to your hard disk in DICOM format. | Availability |
| 2.2.5.19.2 | Export studies and protocols to CD/DVD discs in DICOMDIR format along with a snapshot viewer | Availability |
| 2.2.5.19.3 | Exporting research and protocols to removable USB media | Availability |
| 2.2.5.20 | DICOM support: |  |
| 2.2.5.20.1 | Checking the DICOM connection (Verification as SCU) | Availability |
| 2.2.5.20.2 | Transfer images in DICOM format to an external archive in DICOM format (Storage as SCU) | Availability |
| 2.2.5.20.3 | Print to a DICOM printer (Basic Grayscale Print Management as SCU) | Availability |
| 2.2.5.20.4 | Providing a DICOM compliance document (DICOM compliance Statements) at the client's additional request | Availability |
| 2.2.5.21 | DICOM support: |  |
| 2.2.5.21.1 | Getting a list of research assignments in DICOM | Availability |
| 2.2.5.22 | Functional features of the arm software of a radiologist AWP: |  |
| 2.2.5.22.1 | ability to create up to two protocols for one study with separate storage in the database | Availability |
| 2.2.5.22.2 | availability of templates for creating research protocols with the ability to create new templates | Availability |
| 2.2.5.22.3 | ability to simultaneously work with several doctors from different Departments to create protocols: issue a warning that the study is taken for description | Availability |
| 2.2.5.22.4 | storing all conclusions in a separate reference list of conclusions with the ability to add, edit, or delete conclusions in the reference list without involving technical specialists | Availability |
| 2.2.6 | X-ray power supply unit |  |
| 2.2.6.1 | High-frequency monoblock X-ray power supply unit | Availability |
| 2.2.6.2 | Capacitive energy storage for operation with a 220 V network | Availability |
| 2.2.6.3 | Power consumption from the supply network, kW | Not more than 2,5 |
| 2.2.6.4 | Output power, kW | Not less than 25 |
| 2.2.6.5 | Number of focal spots of the x-ray tube, pcs. | Not less than 2 |
| 2.2.6.6 | Size of the small focal spot of the x-ray tube, mm | Not more than 1,0х1,0 |
| 2.2.6.7 | Size of the big focal spot of the x-ray tube, mm | Not more than 1,5х1.5 |
| 2.2.6.8 | Range of changes in the anode voltage, kV | Not less than 40-125 |
| 2.2.6.9 | Range of changes in the amount of electricity, mA | Not less than 1 – 25 |
| 2.2.6.10 | Ability to work with an unstable supply network in the range from 160 V to 260 V | Availability |
| 2.2.6.11 | Self-diagnosis system of RPU with indication of the error code or malfunction on the display of the remote control | Availability |
| 2.2.7 | Remote control of the X-ray power supply unit |  |
| 2.2.7.1 | Language for indicating and marking controls | Russian |
| 2.2.7.2 | The choice of dimensions of focal spots | Availability |
| 2.2.7.3 | The selection of all exposure parameters (kV, mA, sec.) for a patient of any complexion is performed automatically when the operator presses the high-voltage switch (without trial exposure and without setting any values manually) | Availability |
| 2.2.7.4 | After the image is taken, the exposure parameters (kV, mA, sec.), as well as the values of the effective equivalent radiation dose calculated according to the method of MUC 2.6.1.1797-03, are displayed on the display panel and automatically entered in the patient's account card | Availability |
| 2.2.7.5 | Automatic entry of exposure parameters (kV, mA, sec.), as well as values of the effective equivalent dose in the patient's record card | Availability |
| 2.2.8 | The tripod: stand the transmitter and receiver |  |
| 2.2.8.1 | Type of tripod | Cabless |
| 2.2.8.2 | Type of tripod mount | floor-standing version |
| 2.2.8.3 | Separate lifting racks for the emitter and receiver attached to the floor, without mechanical attachment to each other | Availability |
| 2.2.8.4 | Possibility of free approach of the patient from any side between the emitter and receiver racks | Availability |
| 2.2.8.5 | Range of movement of the emitter and receiver relative to the floor, mm | Not less than 1070 - 1530 |
| 2.2.8.6 | Focal length, mm | Not less than 1000 |
| 2.2.9 | Auxiliary devices and accessories: |  |
| 2.2.9.1 | Gonadal protection skirt with equivalent, Pb | Not less than 0,35 |
| 2.2.9.2 | Protection of the thyroid gland equivalent, Pb | Not less than 0,5 |

Module 2

| № | Item | Item content (value) |
| --- | --- | --- |
| I. | Module 2 | vehicle chassis |
| 1. | General characteristics of Module 2 |  |
| 1.1 | Purpose | For medical assistance during parking |
| 1.2 | Compliance with the requirements of TR CU 018/2011 of the Technical Regulations of the Customs Union. On the safety of wheeled vehicles. | Availability of confirmation in the form of a randomly completed application, certified by the seal of the company. |
| 1.3 | Category of vehicle | N3G |
| 1.4 | Wheel arrangement of the vehicle | 6х6 |
| 1.5 | Axles quantity | Not less than 3 |
| 1.6 | Cabin type | Over the engine location |
| 1.7. | Number of seats in the cabin | Not less than 3 |
| 1.8 | Number of fitted seats in the van designated for location of the passengers on the move | Not less than 7 |
| 1.9 | Engine type | Diesel, four-stroke |
| 1.9.1 | Emission class | Euro-5 |
| 1.9.2 | Cubic capacity of the engine, l | Not less than 6,7 |
| 1.10 | Van type | Multi-section insulated van consisting of sandwich panels with fixed partitions mounted on the vehicle chassis |
| 1.11 | Van dimensions |  |
| 1.11.1 | Length, m | Not less than 5,5 |
| 1.11.2 | Width, m | Not less than 2,4 |
| 1.11.3 | Hight, m | Not less than 2,4 |
| 1.12 | The composition of the exterior sandwich panels | - clad steel coated with white powder enamel or white plastic reinforced with fiberglass, resistant to disinfectant treatment; -plywood; expanded polystyrene; -plywood; - white plastic reinforced with fiberglass, resistant to disinfectant treatment |
| 1.13 | The composition of the inner sandwich panels | - white plastic reinforced with fiberglass, resistant to disinfectant treatment; -plywood; expanded polystyrene; -plywood; - white plastic reinforced with fiberglass, resistant to disinfectant treatment |
| 1.14 | The width of the external sandwich panels, mm | Not less than 60 |
| 1.15 | The width of the inner sandwich panels, mm | Not less than 40 |
| 1.16 | The elements embedded in the sandwich panels for installation of internal and external equipment | Availability |
| 1.17 | Compartments |  |
| 1.17.1 | Medical compartment | with an entry (outdoor) wing door |
| 1.17.2 | Sanitary compartment | with a wing door |
| 1.17.3 | Generator compartment | with a wing door |
| 1.17.4 | Climate control compartment | with a wing door |
| 1.18 | Window | Tinted double-pane glass with light-protective shutters |
| 1.18.1 | Glass panel dimensions | Not less than 700×500 |
| 1.19 | Doors | of sandwich panels with locking devices, holding devices in closed and open position |
| 1.19.1 | The width of the doors sandwich panels, mm | Not less than 40 |
| 1.19.2 | The doors dimensions (HxW), mm | Not less than 1800×600 |
| 1.20 | Floor | unibody construction of sandwich panels laid with linoleum |
| 1.20.1 | The width of the floor sandwich panel, mm | Not less than 80 |
| 1.21 | Accessories (hinges, locks), exterior and interior trim of the van, doors, door openings, joints of the sandwich panels in various plane surfaces | Stainless steel, aluminum extrusion |
| 1.22 | Entry portal | A structure made of aluminium shape with double packing |
| 1.23 | The box beneath the van bottom | Metal welded box with locking device |
| 1.24 | External power supply | Connection to single-phase power grid 220 Volt, m |
| 1.24.1 | The cable length for connection to single-phase power grid 220 Volt, m | Not less than 30 |
| 1.25 | Protective grounding | Availability |
| 1.26 | Autonomous power supply | Gasoline single-phase power supply station, 220 Volt |
| 1.26.1 | The capacity of the gasoline power station, kW | Not less than 9 |
| 1.26.2 | The tank volume of the gasoline power station, l | 10 |
| 1.27 | Autonomous power supply from the on-board power supply system of the vehicle | Via a multifunctional voltage converter with 24 V DC on-board vehicle network to 220 V AC |
| 1.27.1 | Power output of multifunction voltage converter, kW | Not less than 2 |
| 1.28 | Electrical equipment | Input electrical panel with switches for external and autonomous power supply, input voltage control relay, introductory electric automatic circuit breaker and incoming line section of safety shutdown, electrical consumer lines with differential circuit breakers, connectors for external and autonomous power supply. Electrical accessories (sockets, switches) in each compartment |
| 1.29 | The number of individual lines of electrical consumers with differential automatic circuit breakers | Not less than 6 |
| 1.30 | Conditioning system | Split system |
| 1.30.1 | External unit of a split-system | Located in the climate control compartment |
| 1.30.2 | Internal unit of a split system | Located in the medical compartment |
| 1.30.3 | Conditioning system operation modes | cooling, ventilation, heating |
| 1.31 | Ventilation system |  |
| 1.31.1 | Ventilation system of the medical treatment compartment | Supply air filtering and ventilation unit with prepared air distribution |
| 1.31.2 | Ventilation system of the generator compartment | Forced supply and exhaust ventilation system with the capacity sufficient for normal operation of a gasoline power station |
| 1.31.3 | Ventilation system of the climate control compartment | Forced supply and exhaust ventilation system by external unit of the split-system |
| 1.32 | Heating system |  |
| 1.32.1 | Convector heater, pcs. | Not less than 2 |
| 1.32.1.1 | Capacity of the convector heater in the medical compartment, kW | Not less than 1 |
| 1.32.1.2 | Capacity of the convector heater in the sanitary compartment, kW | Not less than 0,5 |
| 1.32.1.3 | Air heat curtain | Above the entry opening in the corridor |
| 1.3.2 | Capacity of the air curtain, kW | Not less than 2 |
| 1.32.2.1 | Autonomous diesel fuel air heater | with the capacity to operate on the move and during parking |
| 1.32.3 | Capacity of the autonomous diesel fuel air heater, kW | Not less than 4 |
| 1.32.3.1 | The basic lighting system | light emitting diode lighting units |
| 1.33 | The emergency lighting system | from the on-board power supply system of the vehicle, 24 V |
| 1.33.1 | Number of emergency lighting unit in each compartment, pcs. | Not less than 1 |
| 1.34 | Fire-fighting appliances | Fire alarm system with the smoke sensors in each compartment; powder-type fire extinguisher |
| 1.34.1 | Capacity of the powder-type fire extinguisher, l | Not less than 2 |
| 1.35 | Sanitary equipment | Availability |
| 1.35.1 | The sink with a tap and a pump for feeding the prepared and heated water in the sanitary compartment | Availability |
| 1.35.2 | Cassette type bio toilet in the sanitary compartment | Availability |
| 1.36 | Furniture | Constructed of panels with laminated working surfaces robust to disinfectant treatment |
| 1.36.1 | Suspended double door shelf in the medical compartment | Availability |
| 1.36.2 | Table in the medical compartment | Availability |
| 1.36.3 | Castor chair with a gas-lift unit in the medical compartment, pcs | Not less than 2 |
| 1.37 | Utility equipment of the compartments | Availability |
| 1.37.1 | Soap dispensers in the sanitary compartment | Availability |
| 1.37.2 | Towel rail in the sanitary compartment | Availability |
| 1.37.3 | Bins, pcs | Not less than 2 |
| 1.37.4 | Clothes hooks, pcs | Not less than 2 |
| 1.37.5 | Container for clean shoe covers, pcs | Not less than 1 |
| 1.37.6 | Container for discarded shoe covers, pcs | Not less than 1 |
| 1.38 | System for fixing medical equipment and other equipment in transport position | Availability |
| 1.39 | Access entry | including the stairs integrated in the van bottom and additional attachable stairs |
| 1.40 | Dashboard camera, pcs. | 1 |
| 1.40.1 | Function | Video and audio record of the situation in front of the vehicle on the move and during parking |
| 1.41 | Subscription telematic server, pcs. | 1 |
| 1.41.1 | Function | Determination and transfer to the telematic server of the location and the movement mode parameters of the vehicle (object) geographical coordinates, speed, course |
| 1.42 | GLONASS/GPS receiver | Availability |
| 1.42.1 | GSM Module | Availability |
| 1.42.2 | Vehicle system of emergency response service, pcs. | 1 |
| 1.42.3 | Repair kit, pcs. | 1 |
| 1.43 | Content of the repair kit | car lifting jack, wheel wrench, warning triangle |
| 2. | Medical equipment |  |
| 2.1 | Dental unit, pcs. | Not less than 1 |
| 2.1.1 | The patient's chair | Availability |
| 2.1.1.1 | Electromechanical drive | Availability |
| 2.1.1.2 | Articulation headrest | Availability |
| 2.1.1.3 | Armrest, pcs. | Not less than 1 |
| 2.1.1.4 | Load capacity of the patient's chair, kg | Not less than 100 |
| 2.1.2 | Doctor's block | Availability |
| 2.1.2.1 | Number of tools | Not less than 3 |
| 2.1.2.2 | Puster-gun water/air | Availability |
| 2.1.2.3 | The micromotor | Availability |
| 2.1.2.4 | Set of tips | Availability |
| 2.1.2.4.1 | Straight tip | Not less than 1 |
| 2.1.2.4.2 | The tip corner | Not less than 1 |
| 2.1.2.4.3 | The tip turbine | Not less than 1 |
| 2.1.2.5 | Polymerization lamp | Availability |
| 2.1.2.6 | Поднос для инструментов | Availability |
| 2.1.3 | Hydroblock | Availability |
| 2.1.3.1 | Type of hydraulic unit | Ejector or vacuum |
| 2.1.3.2 | Type of spittoon | Rotary |
| 2.1.3.3 | Glass filling system | Availability |
| 2.1.3.4 | Spittoon washing system | Availability |
| 2.1.3.5 | Control of chair and instruments | Availability |
| 2.1.4 | Assistant block | Availability |
| 2.1.5 | Saliva ejector | Availability |
| 2.1.6 | Vacuum cleaner | Availability |
| 2.1.7 | Illuminator | Availability |
| 2.1.7.1 | Type of illuminator | Led |
| 2.1.7.2 | Max illumination, Lux | Not less than 30000 |
| 2.1.8 | Dental compressor | Availability |
| 2.1.8.1 | Type of compressor | Oil-free |
| 2.1.8.2 | Sound insulation cover | Availability |
| 2.1.8.3 | Receiver, l | Not less than 10 |
| 2.1.9 | Protective shield against mechanical damage to the eyes for the doctor | Not less than 1 |
| 2.1.10 | Negatoscope portable | Availability |
| 2.2 | X-ray dental device, pcs. | Not less than 1 |
| 2.2.1 | Type of execution | Wall |
| 2.2.2 | Focal spot, mm | Not more than 0,8 |
| 2.2.3 | Maximum automatic shutdown current, A | Not more than 10 |
| 2.2.4 | Remote control | Availability |
| 2.2.5 | Rotating the monoblock 360 degrees horizontally | Availability |
| 2.2.6 | Electronic timer built into the monoblock | Availability |
| 2.2.7 | Weight, kg | Not more than 25 |
| 2.2.8 | Personal protective equipment against x-rays | Availability |
| 2.3 | The intraoral radiology dental, pcs. | Not less than 1 |
| 2.3.1 | Resolution, line pairs / mm | Not less than 20 |
| 2.3.2 | Active surface, mm | Not less than 20×20 |
| 2.3.3 | Cable length, m | Not less than 3 |
| 2.4 | Glasperlen sterilizer, pcs. | Not less than 1 |
| 2.4.1 | Operating temperature in the sterilization chamber, degrees Celsius | Not less than 240 |
| 2.4.2 | Time to enter the temperature stabilization mode, min | Not more than 25 |
| 2.4.3 | Overall dimensions of the sterilizer (LxWxH), mm | Not more than 100×100×110 |
| 2.4.4 | Weight of the sterilizer (without filler), kg | Not more than 1,5 |
| 2.5 | Ultrasonic bath, pcs. | Not less than 1 |
| 2.5.1 | Frequency of the ultrasonic transducer, kHz | Not less than 59 |
| 2.5.2 | Timer, min | Not less than 3 |
| 2.5.3 | Volume of the ultrasonic bath tank, ml | Not less than 150 |
| 2.5.4 | Weight, kg | Not more than 0,5 |
| 2.6 | Dry-burning Cabinet, pcs. | Not less than 1 |
| 2.6.1 | Volume of the sterilization chamber, l | Not less than 10 |
| 2.6.2 | Heating time up to 180 degrees Celsius, min | Not more than 30 |
| 2.6.3 | Emergency disconnection of the sterilizer from the mains in case of overheating in the chamber | Availability |
| 2.6.4 | Continuous operation time per day, h | Not less than 16 |
| 2.6.5 | Overall dimensions (LxWxH), mm | Not more than 450×450×420 |
| 2.7 | UV chamber for preserving sterility, pcs. | Not less than 1 |
| 2.7.1 | Time of continuous operation of the camera, min. | Not less than 160 |
| 2.7.2 | Time for the camera to return to operating mode, min | Not more than 10 |
| 2.7.3 | Automatically turns off the germicidal lamp when opening the lid | Availability |
| 2.7.4 | Overall dimensions (LxWxH), mm | Not more than 600×350×350 |
| 2.7.5 | Weight, kg | Not more than 5 |
| 2.8 | Container for disinfection and pre-sterilization treatment of medical devices, pcs. | Not less than 2 |
| 2.8.1 | Useful volume, l | Not less than 1 |
| 2.8.2 | Total volume, l | Not less than 1,6 |
| 2.8.3 | Weight, kg | Not more than 1 |
| 2.9 | Ultraviolet bactericidal air recirculator, pcs. | Not less than 2 |
| 2.9.1 | Type | Recirculating |
| 2.9.2 | Number of lamps, pcs. | Not less than 3 |
| 2.9.3 | Lamps service life, h | Not less than 9000 |
| 2.9.4 | Power consumption, kW | Not more than 250 |
| 2.9.5 | Performance m3/h | Not less than 50 |
| 2.9.6 | Sterilization efficacy, % | Not less than 99 |
| 2.9.7 | Weight, kg | Not more than 8 |
| 2.10 | Tester the viability of the pulp, pcs. | Not less than 1 |
| 2.10.1 | Display | Availability |
| 2.10.2 | Current range, mkA | Not less than 0 – 80 |
| 2.10.3 | Possibility of battery operation | Availability |
| 2.10.4 | Weight, g | Not more than 100 |
| 2.11 | Apex locator, pcs. | Not less than 1 |
| 2.11.1 | Display | Availability |
| 2.11.2 | Measuring range, mm | Not less than 0 – 3 |
| 2.11.3 | Possibility of battery operation | Availability |
| 2.11.4 | Weight, g | Not more than 50 |
| 2.12 | The electrocoagulator, pcs. | Not less than 1 |
| 2.12.1 | Maximum output power, W | Not less than 5,7 |
| 2.12.2 | Power consumption, W | Not more than 100 |
| 2.12.3 | Overall dimensions (LxWxH), mm | Not more than 235х130х65 |
| 2.13 | Led device | Not less than 1 |
| 2.13.1 | Radiation power, mW | Not more than 400 |
| 2.13.2 | Wavelength, nm | Not less than 395 |
| 2.13.3 | Spectral range, nm | Not less than 470-700 |
| 2.13.4 | Temporary operation mode | Непрерывный |
| 2.13.5 | Power consumption, VA | Not more than 10 |
| 2.13.6 | Battery operation | Availability |
| 2.13.7 | Glasses (red, yellow, green), pcs. | Not less than 3 |
| 2.13.8 | Battery charger | Availability |
| 2.14 | Portable autoclave, pcs. | Not less than 1 |
| 2.14.1 | Does not require connection to a centralized water supply system | Availability |
| 2.14.2 | Drying system | Вакуумная |
| 2.14.3 | Supply system | Встроенный бак для воды |
| 2.14.4 | Built-in water tank capacity, l | Not less than 3 |
| 2.14.5 | The volume of the chamber, l | Not more than 15 |
| 2.14.6 | Display | Availability |
| 2.14.7 | Noise level, dB | Not more than 50 |
| 2.14.8 | Weight, kg | Not more than 50 |
| 2.15 | A set of tools, pcs. | Not less than 1 |
| 2.16 | General practitioner bag with medical diagnostic kits, pcs. | Not less than 1 |
| 2.16.1 | Ophthalmoscope | Availability |
| 2.16.2 | Otoscope with a set of reusable ear funnels | Availability |
| 2.16.3 | Neurological hammer with a brush and a needle | Availability |
| 2.16.4 | Stethoscope complete with two spare headphones and a membrane | Availability |
| 2.16.5 | Peak flow meter | Availability |
| 2.16.6 | Portable blood glucose meter | Availability |
| 2.16.7 | Tonometer | Availability |
| 2.16.8 | Tuning fork 256 Hz | Availability |
| 2.16.9 | Diagnostic illuminator for checking the pupillary reflex, examination of the pharynx and for domestic use | Availability |
| 2.16.10 | Portable electrocardiograph | Availability |
| 2.16.11 | Stacking bag | Availability |
| 2.16.12 | The number of compartments of the bag for equipment, personal belongings of a doctor, documentation and writing instruments | Not less than 3 |
| 2.16.13 | Stacking bag dimensions (L×W×H), mm | Not more than 500×500×260 |

Module 3

| No. | Item | Item content (value) |
| --- | --- | --- |
| I. | Module 3 | on trailer chassis |
| 1. | General specifications of Module 3 |  |
| 1.1 | Purpose | For delivering medical aid during stops |
| 1.2 | Compliance with the requirements of Technical regulations of the Customs Union TRCU 018/2011. On safety of wheeled vehicles. | Availability |
| 1.3 | Possibility of operation within a caravan with Module 1 or Module 2 | Availability |
| 1.4 | Vehicle category | О3 |
| 1.5 | Quantity of axles | At least 2 |
| 1.6 | Van type | Multi-section insulated van made of sandwich panels with fixed partitions installed on the vehicle chassis. |
| 1.7 | Van overall dimensions |  |
| 1.7.1 | Length (m) | At least 7.0 |
| 1.7.2 | Width (m) | At least 2.4 |
| 1.7.3 | Height (m) | At least 2.4 |
| 1.8 | Composition of sandwich panels | - plated steel covered with white powdered enamel or white plastic reinforced with fibre glass resistant to disinfectant treatment; - plywood; - polystyrene foam; - plywood; - white plastic reinforced with fibre glass, resistant to disinfectant treatment |
| 1.9 | Composition of inner sandwich panels | - white plastic reinforced with fibre glass, resistant to disinfectant treatment; - plywood; - polystyrene foam; - plywood; - white plastic reinforced with fibre glass, resistant to disinfectant treatment |
| 1.10 | Thickness of outer sandwich panels (mm) | At least 60 |
| 1.11 | Thickness of inner sandwich panels (mm) | At least 40 |
| 1.12 | Embedded parts in sandwich panels for installation of inner and outer equipment | Availability |
| 1.13 | Compartments |  |
| 1.13.1 | Medical reception compartment No. 1 | with a window and a swing door |
| 1.13.2 | Medical reception compartment No. 2 | with a window and a sliding door |
| 1.13.3 | Medical reception compartment No. 3 | with a window and a sliding door |
| 1.13.4 | Corridor | with an entrance (exterior) swing door |
| 1.13.5 | Sanitary compartment | with a swing door |
| 1.13.6 | Generator compartment | with a swing door |
| 1.13.7 | Air conditioning compartment | with a swing door |
| 1.14 | Windows | Double tinted insulating glass unit with lightproof shades |
| 1.14.1 | Glazing area (mm) | At least 700×500 |
| 1.15 | Doors | made of sandwich panels with locking devices, devices for fixing the door in opened and closed positions |
| 1.15.1 | Door sandwich panel thickness (mm) | At least 40 |
| 1.15.2 | Door overall dimensions (W х T) (mm) | At least 1800×650 |
| 1.16 | Floor | of the load-carrying structure, covered with linoleum |
| 1.16.1 | Floor sandwich panel thickness (mm) | At least 80 |
| 1.17 | Fittings (hinges, lock strikers), exterior and interior finishing of the van, doors, door openings, connections of sandwich panels in different planes | Stainless steel, aluminium shape |
| 1.18 | Entry portal in the corridor | A structure made of aluminium shape with double packing |
| 1.19 | Box under the van bottom | Metal, welded design, with a locking device |
| 1.20 | External power supply | Connection to 220 V single-phase electric mains |
| 1.20.1 | Length of cable for connection to 220 V single-phase electric mains (m) | At least 30 |
| 1.21 | Protective grounding | Availability |
| 1.22 | Self-contained power supply | Petrol-powered electric generating station, 220 V, single-phased |
| 1.22.1 | Petrol-powered electric generation station capacity (kW) | At least 9 |
| 1.22.2 | Petrol-powered electric generation station tank volume (l) | 10 |
| 1.23 | Vehicle self-contained or on-board power supply | Via a multi-functional voltage converter from 24 VDC of the on-board power system to 220 VAC |
| 1.23.1 | Multi-functional voltage converter output power (kW) | At least 2 |
| 1.24 | Electric equipment | Input electric board with switches for external and on-board power supply, input voltage check relay, local circuit breaker and local residual current circuit breaker, electric consumer lines with differential circuit breakers, terminals for connection of external and self-contained power supply. Electric devices (sockets, switches) in each compartment |
| 1.25 | Quantity of separate lines of electric consumers with differential circuit breakers | At least 6 |
| 1.26 | Air conditioning system | Multisplit system |
| 1.26.1 | External unit of the multisplit system | Located in the air conditioning compartment |
| 1.26.2 | Internal units of the multisplit system | Located at least in the medical reception compartments |
| 1.126.3 | Air conditioning system operation modes | cooling, ventilation, heating |
| 1.27 | Ventilation system | Air supply and filtration unit with distribution of treated air to the medical reception compartments |
| 1.27.1 | Generator compartment ventilation system | Forced supply and exhaust ventilation in the volume sufficient for normal operation of petrol-powered electric generation plant |
| 1.27.2 | Air conditioning compartment ventilation system | Forced supply and exhaust ventilation by means of external unit of the multisplit system |
| 1.28 | Heating system | self-contained diesel fuel-powered heater which can operate both on the way and during the stop |
| 1.28.1 | Convector heaters (pcs) | At least 4 |
| 1.28.1.1 | Wattage of convector heater in medical reception compartment No. 1 (kW) | At least 1 |
| 1.28.1.2 | Wattage of convector heater in medical reception compartment No. 2 (kW) | At least 1 |
| 1.28.1.3 | Wattage of convector heater in medical reception compartment No. 3 (kW) | At least 1 |
| 1.28.1.4 | Wattage of convector heater in the sanitary compartment (kW) | At least 0.5 |
| 1.28.2 | Air heat curtain | In the corridor, above the entrance opening |
| 1.28.2.1 | Wattage of the air heat curtain (kW) | At least 2 |
| 1.28.3 | Self-contained diesel fuel-powered air heater | which can operate on the way and during the stop (parking) |
| 1.28.3.1 | Wattage of self-contained diesel fuel-powered air heater (kW) | At least 4 |
| 1.29 | The main illumination system | LED light fixtures |
| 1.30 | The emergency illumination system | powered by 24 V vehicle on-board system |
| 1.30.1 | Quantity of emergency light fixtures in each compartment (pcs) | At least 1 |
| 1.31 | Fire-fighting equipment | Fire alarm system with smoke detectors in each compartment; dry powder extinguisher |
| 1.31.1 | Dry powder extinguisher volume (l) | At least 2 |
| 1.32 | Sanitary facilities | Availability |
| 1.32.1 | Washer with a tap and pump for supply of treated heated water in the medical reception compartments and in sanitary compartment | Availability |
| 1.32.2 | Cassette-type portable biotoilet in the sanitary section | Availability |
| 1.33 | Furniture | Made of panels with lamination of operating surfaces resistant to disinfectant treatment. |
| 1.33.1 | Suspended two-fold shelves in the medical reception compartments | Availability |
| 1.33.2 | Table, pcs | At least 3 |
| 1.33.3 | Castor chair with a gas-lift unit (pcs) | At least 3 |
| 1.33.4 | Castor stool with a gas-lift unit (pcs) | At least 3 |
| 1.33.5 | Stool in the corridor | At least 3 |
| 1.34 | Utility equipment of the compartments | Availability |
| 1.34.1 | Soap dispensers in the medical reception compartments and in the sanitary compartment | Availability |
| 1.34.2 | Towel holders in the medical reception compartments and in the sanitary compartment | Availability |
| 1.34.3 | Garbage cans (pcs) | At least 4 |
| 1.34.4 | Clothes hooks (pcs) | At least 4 |
| 1.34.5 | Container for clean shoe covers (pcs) | At least 1 |
| 1.34.6 | Container for used shoe covers (pcs) | At least 1 |
| 1.35 | A system for fixing medical equipment and other equipment in transport position | Availability |
| 1.36 | Entry portal | with steps integrated into the van bottom and additional removable steps |
| 2. | Medical equipment: |  |
| 2.1. | Gynecological examination chair (pcs) | At least 1 |
| 2.1.1. | Electrical adjustment of height, seat angle and seatback angle | Availability |
| 2.1.2. | Pan | Availability |
| 2.1.3. | Geppel leg supports | Availability |
| 2.1.4. | Transformability into bed | Availability |
| 2.1.5. | Leg section | Availability |
| 2.1.6. | Maximum load (kg) | At least 170 |
| 2.2 | Wall-mounted lighting fixture | At least 2 |
| 2.2.1 | Flexible holder | Availability |
| 2.2.2 | LED lamp | Availability |
| 2.2.3 | Handle with rheostat | Availability |
| 2.2.4 | Hardware version | Wall-mounted |
| 2.3. | Colposcope on a console support | At least 1 |
| 2.3.1 | Microscope type | Binocular |
| 2.3.2 | LED illumination | Availability |
| 2.3.3 | Q-ty of magnification steps | At least 5 |
| 2.3.4 | Cyan filter | Availability |
| 2.4. | Microscope (pcs) | At least 1 |
| 2.4.1. | Design type | Binocular or monocular |
| 2.5. | Portable ultrasonic system with accessories | At least 1 |
| 2.5.1. | Possible application fields | Minimum: Abdominal examinations Obstetrics and gynaecology Neurology Traumatology and orthopedics Oncology Endocrinlogy Pediatrics Neonatology Urology |
| 2.5.2. | Quantity of positions in terms of focus area scanning depth for radiation | at least 4 |
| 2.5.3. | Real-time image magnification (acoustic zoom) | at least 10 |
| 2.5.4. | Image magnification in still-frame mode (PAN zoom) | at least 10 |
| 2.5.5. | Full-screen canning mode for only ultrasonic image activated by pressing a special button | Availability |
| 2.5.6. | Maximum system dynamic range value displayed on the screen (dB) | at least 220 |
| 2.5.7. | System frequency range (MHz) | at least 2–14.0 |
| 2.5.8. | Quantity of digital processing channels | at least 1,024 |
| 2.5.9. | Maximum frame rate per second | at least 400 |
| 2.5.10. | Quantity of ultrasonic radiation frequencies for scanning B mode | at least 4 |
| 2.5.11. | Quantity of tissue harmonic frequencies | at least 2 |
| 2.5.12. | Multipath reception | Availability |
| 2.5.13. | Tissue inverse harmonics with phase shift | Availability |
| 2.5.14. | Cine loop, quantity of frames. | at least 10,500 |
| 2.5.15. | Reporting in Russian with possibility to add images and comments | Availability |
| 2.5.16. | Saving reports in PDF/RTF format | Availability |
| 2.5.17. | Archiving the image to integrated hard disk drive | Availability |
| 2.5.18. | Archiving the image to CD/DVD drives (if an external disk drive is available) | Availability |
| 2.5.19. | Archiving the image to external media via USB port | Availability |
| 2.5.20. | Memory size for information storage, GB | at least 500 |
| 2.5.21. | Colour monitor | Availability |
| 2.5.22. | Monitor diagonal size (inches) | at least 15 |
| 2.5.23. | Monitor resolution (pixels) | at least 1024х768 |
| 2.5.24. | Russified control panel | Availability |
| 2.5.25. | Built-in keyboard | Availability |
| 2.5.26. | B mode | Availability |
| 2.5.27. | M mode | Availability |
| 2.5.28. | Pseudoconvex scanning in B mode for linear array probes | Availability |
| 2.5.29. | Spatial compound imaging | Availability |
| 2.5.30. | Mode of the second tissue harmonics | Availability |
| 2.5.31. | Adjustment of acoustic radiation power with display of values on the monitor screen | Availability |
| 2.5.32. | Adjustment of the received signal amplification with display of values on the monitor screen | Availability |
| 2.5.33. | Dynamic focusing for reception | Availability |
| 2.5.34. | Dynamic aperture for transmission and reception | Availability |
| 2.5.35. | Adjustability of line density in B mode | Availability |
| 2.5.36. | Dynamic filtration in terms of scanning depth | Availability |
| 2.5.37. | Pseudocolouring of halftone image | Availability |
| 2.5.38. | Changing the image parameters on still image | Availability |
| 2.5.39. | Image rotation and inversion | Availability |
| 2.5.40. | Image boundary highlighting filter | Availability |
| 2.5.41. | Image smoothing | Availability |
| 2.5.42. | Gamma correction control | Availability |
| 2.5.43. | Cine loop playback rate control | Availability |
| 2.5.44. | Archiving prospective and retrospective cine loop with arbitrary setting of its length | Availability |
| 2.5.45. | Cine loop editability, including segmental editing | Availability |
| 2.5.46. | Convex probe | Availability |
| 2.5.47. | Linear array probe | Availability |
| 2.5.48. | Intracavitary probe | Availability |
| 2.5.49. | Printer | Availability |
| 2.5.50. | Special cart | Availability |
| 2.5.51. | Battery | Availability |
| 2.5.52. | Overall dimensions of the main unit (LхWхH) (mm) | not more than 500 х 300 х 400 |
| 2.5.53. | Weight of the main unit (kg) | not more than 10 |
| 2.6. | Manual artificial respirating unit (pcs) | At least 1 |
| 2.6.1. | Breathing bag for manual ALV (pcs) | At least 1 |
| 2.6.2. | Tongue depressor (pcs) | 1 |
| 2.6.3. | Mouth gag (pcs) | 1 |
| 2.6.4. | Carrier bag (pcs) | Availability |
| 2.7 | A general kit for delivery of first aid (pcs) | At least 1 |
| 2.8 | Electrocardiograph (pcs) | 1 |
| 2.8.1. | Recording, q-ty of channels | At least 12 |
| 2.8.2. | Paper feed rate (mm/s) | 5; 10; 25; 50 |
| 2.8.3. | Device memory, quantity of examinations | At least 300 |
| 2.8.4. | Display resolution (pixels) | At least 640×480 |
| 2.8.5. | Protection against defibrillation | Availability |
| 2.8.6. | Battery operation time (h) | At least 3 |
| 2.8.7. | Menu control using a rotating manipulator | Availability |
| 2.8.8. | Thermal paper width (mm) | At least 110 |
| 2.8.9. | Overall dimensions |  |
| 2.8.10. | Length (mm) | Not more than 300 |
| 2.8.11. | Width (mm) | Not more than 200 |
| 2.8.12. | Height, mm | Not more than 100 |
| 2.8.13. | Weight with battery and patient cable (kg) | Not more than 4 |
| 2.9. | Diagnostic spirometer (pcs) | At least 1 |
| 2.9.1. | Spirometery tests | At least: FVC, VC |
| 2.9.2. | Display | Availability |
| 2.9.3. | Software | Availability |
| 2.9.4. | Sensor | turbine |
| 2.9.5. | Communication port | Availability |
| 2.9.6. | Weight (kg) | Not more than 0.4 |
| 2.10. | Stethoscope (pcs) | At least 2 |
| 2.10.1. | Double-sided membrane | Availability |
| 2.10.2. | Soft replaceable caps for earphones | Availability |
| 2.10.3. | PVC or equivalent Y-tubes | Availability |
| 2.10.4. | Total length (cm) | At least 55 |
| 2.11. | Cardiological stethoscope (pcs) | At least 1 |
| 2.11.1. | Double-sided membrane | Availability |
| 2.11.2. | Soft replaceable caps for earphones | Availability |
| 2.11.3. | PVC or equivalent Y-tubes | Availability |
| 2.11.4. | Total length (cm) | At least 74 |
| 2.12. | Wall-mounted sphygmomanometer with a dial | At least 3 |
| 2.12.1. | Overall dimensions (mm) | At least 150×150 |
| 2.12.2. | Insufflation tube | Availability |
| 2.12.3. | Wall mounting | Availability |
| 2.13. | Height meter | At least 1 |
| 2.13.1. | Design type | Wall-mounted |
| 2.13.1. | Measurement range (m) | No less than 0.8–2.0 |
| 2.14. | Medical balance (pcs) | At least 1 |
| 2.14.1. | Minimum measured weight (kg) | At least 150 |
| 2.14.2. | Display | Availability |
| 2.15. | Pedobarometer (pcs) | At least 1 |
| 2.15.1. | Weighing limit (kg) | At least 15 |
| 2.15.2. | Reading setting time (s) | Not more than 2 |
| 2.15.3. | Overall dimensions (including the removable platform) (LхWхH) (mm) | Not more than 540×330×105 |
| 2.15.4. | Overall dimensions of the removable platform (LхWхH) (mm) | Not more than 540×290×70 |
| 2.15.5. | Weight (kg) | Not more than 5 |
| 2.16. | Blood analyzer biochemical semi-automatic (pcs) | Not more than 1 |
| 2.16.1. | Light source | halogen lamp |
| 2.16.2. | The wavelength range, nm | No less than 340-500 |
| 2.16.3. | Filters, pcs | No less than 5 |
| 2.16.4. | Display | Availability |
| 2.16.5. | Integrated printer | Availability |
| 2.16.6. | Measurement error, % | Not more than 5 |
| 2.16.7. | Weight, kg | Not more than 25 |
| 2.17. | Glucometer (pcs) | At least 1 |
| 2.17.1. | Specimen | Fresh capillary blood |
| 2.17.2. | Blood drop volume (mmol/l) | Not more than 0.5 |
| 2.17.3. | Measurement duration (s) | Not more than 5 |
| 2.17.4. | Range of measured values (mmol/l) | No less than 1.1–33.3 |
| 2.17.5. | Memory size, quantity of results | At least 100 |
| 2.17.6. | Automatic switch-off | Availability |
| 2.17.7. | Overall dimensions (LхWхH) (cm) | Not more than 10×10×2 |
| 2.17.8. | Weight (kg) | Not more than 0.2 |
| 2.18. | Portable urine analyser (pcs) | At least 1 |
| 2.18.1. | Quantity of measured tests | At least 10 |
| 2.18.2. | Performance (tests per hour) | At least 60 |
| 2.18.3. | Built-in thermal printer | Availability |
| 2.18.4. | Device memory (quantity of examinations) | At least 100 |
| 2.18.5. | Dimensions (mm) | Not more than 300 х 200 х 200 |
| 2.18.6. | Weight (kg) | Not more than 2 |
| 2.19. | Stand (pcs) | At least 2 |
| 2.19.1. | Quantity of sockets | At least 50 |
| 2.20. | Container for disinfection and presterilising treatment of medical devices (pcs) | At least 4 |
| 2.20.1. | Exterior overall dimensions (mm) | Not more than 250×170×100 |
| 2.20.2. | Effective volume (l) | At least 1 |
| 2.20.3. | Full volume (l) | At least 1.6 |
| 2.20.4. | Weight (kg) | Not more than 0.7 |
| 2.21. | Ultraviolet bactericidal air recirculator (pcs) | At least 3 |
| 2.21.1. | Type | Recirculating |
| 2.21.2. | Quantity of lamps (pcs) | At least 3 |
| 2.21.3. | Lamp service life (h) | At least 9000 |
| 2.21.4. | Power consumption (W) | Not more than 250 |
| 2.21.5. | Delivery (m3/h) | At least 50 |
| 2.21.6. | Bactericidal efficiency (%) | At least 99 |
| 2.21.7. | Weight (kg) | Not more than 8 |
| 2.22. | Blood analyzer hematological (pcs) | No less than 1 |
| 2.22.1. | Number of parameters to define | No less than 18 |
| 2.22.2. | Volume of the whole blood sample (µсl) | Not more than 10 |
| 2.22.3. | The volume of sample dilution (µсl) | 20 |
| 2.22.4. | Productivity, analyses per hour | No less than 40 |
| 2.22.5. | Built-in thermal printer | Availability |
| 2.22.6. | Dimensions (mm) | Not more than 500 х 500 х 400 |
| 2.22.7. | Weight (kg) | Not more than 20 |

Module 4

| No. | Item | Item content (value) |
| --- | --- | --- |
| I. | Module 4 | on trailer chassis |
| 1. | General specifications of Module 4 |  |
| 1.1 | Purpose | For administering medical aid during stops |
| 1.2 | Compliance with the requirements of Technical regulations of the Customs Union TRCU 018/2011. On safety of wheeled vehicles. | Availability |
| 1.3 | Possibility of operation within a caravan with Module 1 or Module 2 | Availability |
| 1.4 | Vehicle category | О3 |
| 1.5 | Quantity of axles | At least 2 |
| 1.6 | Van type | Multi-section insulated van made of sandwich panels with fixed partitions installed on the vehicle chassis. |
| 1.7 | Van overall dimensions |  |
| 1.7.1 | Length (m) | At least 7.0 |
| 1.7.2 | Width (m) | At least 2.4 |
| 1.7.3 | Height (m) | At least 2.4 |
| 1.8 | Composition of sandwich panels | - plated steel covered with white powdered enamel or white plastic reinforced with fibre glass resistant to disinfectant treatment; - plywood; - polystyrene foam; - plywood; - white plastic reinforced with fibre glass, resistant to disinfectant treatment |
| 1.9 | Composition of inner sandwich panels | - white plastic reinforced with fibre glass, resistant to disinfectant treatment; - plywood; - polystyrene foam; - plywood; - white plastic reinforced with fibre glass, resistant to disinfectant treatment |
| 1.10 | Thickness of outer sandwich panels (mm) | At least 60 |
| 1.11 | Thickness of inner sandwich panels (mm) | At least 40 |
| 1.12 | Embedded parts in sandwich panels for installation of inner and outer equipment | Availability |
| 1.13 | Compartments |  |
| 1.13.1 | Medical reception compartment No. 1 | with a window and a swing door |
| 1.13.2 | Medical reception compartment No. 2 | with a window and a sliding door |
| 1.13.3 | Medical reception compartment No. 3 | with a window and a sliding door |
| 1.13.4 | Corridor | with an entrance (exterior) swing door |
| 1.13.5 | Sanitary compartment | with a swing door |
| 1.13.6 | Generator compartment | with a swing door |
| 1.13.7 | Air conditioning compartment | with a swing door |
| 1.14 | Windows | Double tinted insulating glass unit with lightproof shades |
| 1.14.1 | Glazing area (mm) | At least 700×500 |
| 1.15 | Doors | made of sandwich panels with locking devices, devices for fixing the door in opened and closed positions |
| 1.15.1 | Door sandwich panel thickness (mm) | At least 40 |
| 1.15.2 | Door overall dimensions (W х T) (mm) | At least 1800×650 |
| 1.16 | Floor | of the load-carrying structure, covered with linoleum |
| 1.16.1 | Floor sandwich panel thickness (mm) | At least 80 |
| 1.17 | Fittings (hinges, lock strikers), exterior and interior finishing of the van, doors, door openings, connections of sandwich panels in different planes | Stainless steel, aluminium shape |
| 1.18 | Entry portal in the corridor | A structure made of aluminium shape with double packing |
| 1.19 | Box under the van bottom | Metal, welded design, with a locking device |
| 1.20 | External power supply | Connection to 220 V single-phase electric mains |
| 1.20.1 | Length of cable for connection to 220 V single-phase electric mains (m) | At least 30 |
| 1.21 | Protective grounding | Availability |
| 1.22 | Self-contained power supply | Petrol-powered electric generating station, 220 V, single-phased |
| 1.22.1 | Petrol-powered electric generation station capacity (kW) | At least 9 |
| 1.22.2 | Petrol-powered electric generation station tank volume (l) | 10 |
| 1.23 | Vehicle self-contained or on-board power supply | Via a multi-functional voltage converter from 24 VDC of the on-board power system to 220 VAC |
| 1.23.1 | Multi-functional voltage converter output power capacity (kW) | At least 2 |
| 1.24 | Electric equipment | Input electric board with switches for external and on-board power supply, input voltage check relay, local circuit breaker and local residual current circuit breaker, electric consumer lines with differential circuit breakers, terminals for connection of external and self-contained power supply. Electric devices (sockets, switches) in each compartment |
| 1.25 | Quantity of separate lines of electric consumers with differential circuit breakers | At least 6 |
| 1.26 | Air conditioning system | Multisplit system |
| 1.26.1 | External unit of the multisplit system | Located in the air conditioning compartment |
| 1.26.2 | Internal units of the multisplit system | Located at least in the medical reception compartments |
| 1.126.3 | Air conditioning system operation modes | cooling, ventilation, heating |
| 1.27 | Ventilation system | Air supply and filtration unit with distribution of treated air to the medical reception compartments |
| 1.27.1 | Generator compartment ventilation system | Forced supply and exhaust ventilation in the volume sufficient for normal operation of petrol-powered electric generation plant |
| 1.27.2 | Air conditioning compartment ventilation system | Forced supply and exhaust ventilation by means of external unit of the multisplit system |
| 1.28 | Heating system | self-contained diesel fuel-powered heater which can operate both on the way and during the stop |
| 1.28.1 | Convector heaters (pcs) | At least 4 |
| 1.28.1.1 | Wattage of convector heater in medical reception compartment No. 1 (kW) | At least 1 |
| 1.28.1.2 | Wattage of convector heater in medical reception compartment No. 2 (kW) | At least 1 |
| 1.28.1.3 | Wattage of convector heater in medical reception compartment No. 3 (kW) | At least 1 |
| 1.28.1.4 | Wattage of convector heater in the sanitary compartment (kW) | At least 0.5 |
| 1.28.2 | Air heat curtain | In the corridor, above the entrance opening |
| 1.28.2.1 | Wattage of the air heat curtain (kW) | At least 2 |
| 1.28.3 | Self-contained diesel fuel-powered air heater | which can operate on the way and during the stop |
| 1.28.3.1 | Wattage of self-contained diesel fuel-powered air heater (kW) | At least 4 |
| 1.29 | The main illumination system | LED light fixtures |
| 1.30 | The emergency illumination system | powered by 24 V vehicle on-board system |
| 1.30.1 | Quantity of emergency light fixtures in each compartment (pcs) | At least 1 |
| 1.31 | Fire-fighting equipment | Fire alarm system with smoke detectors in each compartment; dry powder extinguisher |
| 1.31.1 | Dry powder extinguisher volume (l) | At least 2 |
| 1.32 | Sanitary facilities | Availability |
| 1.32.1 | Washer with a tap and pump for supply of treated heated water in the medical reception compartments and in sanitary compartment | Availability |
| 1.32.2 | Cassette-type portable toilet | Availability |
| 1.33 | Furniture | Made of panels with lamination of operating surfaces resistant to disinfectant treatment. |
| 1.33.1 | Suspended two-fold shelves in the medical reception compartments | Availability |
| 1.33.2 | Table, pcs | At least 3 |
| 1.33.3 | Castor chair with a gas-lift unit (pcs) | At least 3 |
| 1.33.4 | Castor stool with a gas-lift unit (pcs) | At least 3 |
| 1.33.5 | Stool in the corridor | At least 3 |
| 1.34 | Utility equipment of the compartments | Availability |
| 1.34.1 | Soap dispensers in the medical reception compartments and in the sanitary compartment | Availability |
| 1.34.2 | Towel holders in the medical reception compartments and in the sanitary compartment | Availability |
| 1.34.3 | Garbage cans (pcs) | At least 4 |
| 1.34.4 | Clothes hooks (pcs) | At least 4 |
| 1.34.5 | Container for clean shoe covers (pcs) | At least 1 |
| 1.34.6 | Container for used shoe covers (pcs) | At least 1 |
| 1.35 | A system for fixing medical equipment and other equipment in transport position | Availability |
| 1.36 | Entry portal | with steps integrated into the van bottom and additional removable steps |
| 2. | Medical equipment: |  |
| 2.1. | Ophthalmoscope with accessories (pcs) | At least 1 |
| 2.1.1. | Corrective lenses | At least 20 |
| 2.1.2. | Apertures | At least 3 |
| 2.1.3. | Wall-mounted module with two handles | Availability |
| 2.1.4. | Additional module (dispenser) | Availability |
| 2.2. | Otoscope with accessories (pcs) | At least 1 |
| 2.2.1. | View window with magnification | Availability |
| 2.2.2. | Possibility of operation from the wall-mounted module, voltage | Availability |
| 2.2.3. | Funnel dispenser | Availability |
| 2.3. | A set of test spectacle lenses with a frame (pcs) | At least 1 |
| 2.3.1. | Quantity of set elements (pcs) | At least 96 |
| 2.3.2. | Stigmatic plus lenses with refraction | Availability |
| 2.3.3. | Stigmatic plus lenses with refraction | Availability |
| 2.3.4. | Stigmatic minus lenses with refraction | Availability |
| 2.3.5. | Stigmatic minus lenses with refraction | Availability |
| 2.3.6. | Astigmatic minus lenses with refraction | Availability |
| 2.3.7. | Astigmatic minus lenses with refraction | Availability |
| 2.3.8. | Astigmatic plus lenses with refraction | Availability |
| 2.3.9. | Astigmatic plus lenses with refraction | Availability |
| 2.3.10. | Red filter | Availability |
| 2.3.11. | Green filter | Availability |
| 2.3.12. | Diaphragm | Availability |
| 2.3.13. | Occluder (screen) | Availability |
| 2.3.14. | Test frame | Availability |
| 2.3.15. | Case |  |
| 2.4. | Light for visual charts (pcs) | At least 1 |
| 2.4.1. | Possibility of applying in a room at a distance of 2.7 m | Availability |
| 2.4.2. | Illumination intensity in the centre of the chart (lux) | At least 700 |
| 2.4.3. | Power consumption (W) | Not more than 100 |
| 2.4.4. | Light | Availability |
| 2.4.5. | A set of charts, including those for children | Availability |
| 2.4.6. | Baton | Availability |
| 2.4.7. | Occluder (screen) | Availability |
| 2.4.8. | Overall dimensions (LxHxW) (mm) | Not more than 335 х 485 х 125 |
| 2.4.9. | Weight (kg) | Not more than 10 |
| 2.5. | Air-puff tonometer for ophthalmotonous pressure measurement (pcs) | At least 1 |
| 2.5.1. | Duration of one measurement (s) | not more than 5 |
| 2.5.2. | Quantity of measurements performed with one set of batteries | At least 1000 |
| 2.5.3. | Overall dimensions (mm) | No more than 200х50х50 |
| 2.5.4. | Weight of the device with batteries (kg) | not more than 0.15 |
| 2.6. | Manual portable slit lamp (pcs) |  |
| 2.6.1. | Microscope | Availability |
| 2.6.2. | Magnification factor | At least 5 |
| 2.6.3. | Light | Availability |
| 2.6.4. | Light slit width change range (mm) | At least 0 to 11 |
| 2.6.5. | Light slit height (mm) | At least 11 |
| 2.6.6. | Built-in filters | Cobalt blue, green, conversion |
| 2.6.7. | Light source | LED |
| 2.6.8. | Battery (pcs) | Availability |
| 2.6.9. | Battery charger | Availability |
| 2.7. | Stethoscope (pcs) | At least 2 |
| 2.7.1. | Double-sided membrane | Availability |
| 2.7.2. | Soft replaceable caps for earphones | Availability |
| 2.7.3. | PVC or equivalent Y-tubes | Availability |
| 2.7.4. | Total length (cm) | At least 55 |
| 2.8. | Wall-mounted sphygmomanometer with a dial | At least 3 |
| 2.8.1. | Overall dimensions (mm) | At least 150×150 |
| 2.8.2. | Insufflation tube | Availability |
| 2.8.3. | Wall mounting | Availability |
| 2.9. | Ultraviolet bactericidal air recirculator (pcs) | 4 |
| 2.9.1. | Type | Recirculating |
| 2.9.2. | Quantity of lamps (pcs) | At least 3 |
| 2.9.3. | Lamp service life (h) | At least 9000 |
| 2.9.4. | Power consumption (W) | Not more than 250 |
| 2.9.5. | Delivery (m3/h) | At least 50 |
| 2.9.6. | Bactericidal efficiency (%) | At least 99 |
| 2.9.7. | Weight, kg | Not more than 8 |

### 4.Drawings

These Bidding Documents includes NO drawings.

[If documents shall be included, insert the following List of Drawings]

| List of Drawings |  |  |
| --- | --- | --- |
| Drawing Nr. | Drawing Name | Purpose |

5.Inspections and Tests

The following inspections and tests shall be performed: NOT APPLICABLE